## Slide 1
Charly Careddaa , Frédéric Langeb , Niccole Ranaei-Zamanic , Uzair Hakimb ,
Olayinka Kowobaric , Dimitrios Siassakosc , Sara Hillmanc , Anna L Davidc , Subhabrata Mitrac and Ilias Tachtsidisb
a Univ Lyon, INSA‐Lyon, Université Claude Bernard Lyon 1, UJM-Saint Etienne, CNRS, Inserm, CREATIS UMR 5220, U1294, Lyon, France
b Department of Medical Physics and Biomedical Engineering, University College London, London, UK
c EGA Institute for Women’s Health, University College London, London, United Kingdom
June 25, 2025
Optimization of continuous-wave NIRS devices for placental monitoring. A simulation study

## Slide 2
Placental function monitoring
1
Abnormal placenta development is a cause of adverse pregnancy outcomes [1]
Measuring placental oxygenation can be used to monitor the placental function
Clinical ultrasound devices do not include direct functional assessment of the placenta
MRI techniques are not suited for bedside monitoring
[1] G. J. Burton et al. Physiol Rev. 2016 Oct;96(4):1509-65 [2] Image taken from https://www.abclawcenters.com/practice-areas/prenatal-birth-injuries/labor-and-delivery-complications-and-errors/placenta-previa/
Fetal ultrasound exam [2]

## Slide 3
Problematic – Placenta sensitivity
What is the sensitivity of CW NIRS devices for monitoring placental function?
What is the best experimental configuration to monitor the placenta?
<number>
Source detector separation (cm)
3
4
5
Mini CYRIL [3]
In-house developed broadband NIRS device
Maternal abdomen monitoring with in-house
developed broadband NIRS device at UCLH
[3] Kaynezhad, P. et al., Neurophotonics 6(4), 045009–045009 (2019).

## Slide 4
Problematic – Placenta sensitivity
What is the sensitivity of CW NIRS devices for monitoring placental function?
What is the best experimental configuration to monitor the placenta?
The measurement of placenta is not direct (skin, adipose tissue, muscle layers)
Emitted light source
Placenta
Muscle
Adipose tissue
Skin
Collected light for detectors at
3 cm
4 cm
<number>
Source detector separation (cm)
3
4
5
Mini CYRIL [3]
In-house developed broadband NIRS device
[3] Kaynezhad, P. et al., Neurophotonics 6(4), 045009–045009 (2019).
Ultrasound image of the maternal abdomen
collected by the clinical team at UCLH

## Slide 5
Study the light propagation in the maternal abdomen with Monte Carlo simulations [4]
<number>
[4] Fang Q and Boas D (2009) Opt. Express, vol. 17, issue 22, pp. 20178-20190
Light source
109 packets of photons
780 nm
One simulation per configuration
NVIDIA GeForce RTX 4080
51 840 simulations
Computation time ≈ 5 days (8 s per simulation)
Output
Diffuse reflectance (1.mm-2.s-1)
Fluence rate (1.mm-2.s-1)
Placenta sensitivity – Monte Carlo simulations
Four layers Monte Carlo model
Source detector
separation (cm)
5
4
3
 | Skin | Adipose tissue | Muscle | Placenta
Thickness (mm) | [1, 2, 3, 4, 5] | [1, 2, 4, 8, 16] | [3, 7, 10, 13, 20, 25] | Semi-infinite
Melanin volume fraction (%) | [2.55, 15.5, 30.5] | - | - | -
Oxygen saturation (%) | 39 (dermis) | - | [40, 60, 80] | [40, 60, 80]
Blood volume (µM) | 4,7 (dermis) | - | [15, 25, 35, 50] | [15, 25, 35, 50]
Water volume fraction (%) | 0 | 80 | 76 | 85
Fat volume fraction (%) | 0 | 20 | 0 | 0
Scattering coefficient (cm-1) |  |  |  | 
Refractive index n | 1.47 | 1.37 | 1.37 | 1.4
Anisotropy coefficient g | 0.83 | 0.9 | 0.9 | 0.9

## Slide 6
Study the light propagation in the maternal abdomen with Monte Carlo simulations [4]
<number>
[4] Fang Q and Boas D (2009) Opt. Express, vol. 17, issue 22, pp. 20178-20190
Light source
109 packets of photons
780 nm
One simulation per configuration
NVIDIA GeForce RTX 4080
51 840 simulations
Computation time ≈ 5 days (8 s per simulation)
Output
Diffuse reflectance (1.mm-2.s-1)
Fluence rate (1.mm-2.s-1)
Placenta sensitivity – Monte Carlo simulations
Four layers Monte Carlo model
Source detector
separation (cm)
5
4
3
 | Skin | Adipose tissue | Muscle | Placenta
Thickness (mm) | [1, 2, 3, 4, 5] | [1, 2, 4, 8, 16] | [3, 7, 10, 13, 20, 25] | Semi-infinite
Melanin volume fraction (%) | [2.55, 15.5, 30.5] | - | - | -
Oxygen saturation (%) | 39 (dermis) | - | [40, 60, 80] | [40, 60, 80]
Blood volume (µM) | 4,7 (dermis) | - | [15, 25, 35, 50] | [15, 25, 35, 50]
Water volume fraction (%) | 0 | 80 | 76 | 85
Fat volume fraction (%) | 0 | 20 | 0 | 0
Scattering coefficient (cm-1) |  |  |  | 
Refractive index n | 1.47 | 1.37 | 1.37 | 1.4
Anisotropy coefficient g | 0.83 | 0.9 | 0.9 | 0.9

## Slide 7
Tissue thickness with ultrasound images
Skin tones with Fitzpatrick scale [5]
Measurements performed at UCLH on 142 healthy participants
<number>
Tissue thickness, placenta depth and skin tone distributions
[5] Karsten, A. E. et al. (2012), Photochemistry and photobiology 88(2), 469–474 (2012).
Clinical measurements
Thickness percentiles (%) | Placenta depth (mm) | Skin thickness (mm) | Adipose tissue thickness (mm) | Muscle thickness (mm)
25 | 10 | 1 | 2 | 7
50 | 16 | 2 | 4 | 10
75 | 20 | 3 | 5 | 12
Ultrasound image of the maternal abdomen
collected by the clinical team at UCLH

## Slide 8
Tissue thickness with ultrasound images
Skin tones with Fitzpatrick scale [5]
Measurements performed at UCLH on 142 healthy participants
<number>
Tissue thickness, placenta depth and skin tone distributions
[5] Karsten, A. E. et al. (2012), Photochemistry and photobiology 88(2), 469–474 (2012).
Clinical measurements
Thickness percentiles (%) | Placenta depth (mm) | Skin thickness (mm) | Adipose tissue thickness (mm) | Muscle thickness (mm)
25 | 10 | 1 | 2 | 7
50 | 16 | 2 | 4 | 10
75 | 20 | 3 | 5 | 12
Ultrasound image of the maternal abdomen
collected by the clinical team at UCLH

## Slide 9
Probability that the detected light at detector D has passed through the placenta
Calculation of the sensitivity profile with the adjoint method [6]:
Voxel-wise multiplication between FS et FD (Fluence rate for source at position S and D)
Normalization of the sensitivity profile by its sum (probability density P)
<number>
Placenta sensitivity
[6] Yao, R et al. Biomedical optics express 9(10), 4588–4603 (2018)

## Slide 10
Probability that the detected light at detector D has passed through the placenta
<number>
Placenta sensitivity
Emitted light source S
Placenta
Muscle
Adipose tissue
Skin
Collected light for detectors at
D cm
J cm
z1
z2
Placenta sensitivity
z

## Slide 11
Probability that the light has passed through the placenta
<number>
Placenta sensitivity versus blood volume and SD-separation
Muscle blood volume: 25 µMol
Muscle SatO2 = 60% Placenta SatO2 = 80%
Thickness percentiles (%) | Placenta depth (mm) | Skin thickness (mm) | Adipose tissue thickness (mm) | Muscle thickness (mm)
25 | 10 | 1 | 2 | 7

## Slide 12
Probability that the light has passed through the placenta
<number>
Placenta sensitivity versus blood volume and SD-separation
Muscle blood volume: 25 µMol
Muscle SatO2 = 60% Placenta SatO2 = 80%
Thickness percentiles (%) | Placenta depth (mm) | Skin thickness (mm) | Adipose tissue thickness (mm) | Muscle thickness (mm)
25 | 10 | 1 | 2 | 7
Placental sensitivity decreases with increasing placental blood volume

## Slide 13
Probability that the light has passed through the placenta
<number>
Placenta sensitivity versus blood volume and SD-separation
Muscle blood volume: 25 µMol
Muscle SatO2 = 60% Placenta SatO2 = 80%
Thickness percentiles (%) | Placenta depth (mm) | Skin thickness (mm) | Adipose tissue thickness (mm) | Muscle thickness (mm)
25 | 10 | 1 | 2 | 7
Placental sensitivity increases with increasing source-detector separation

## Slide 14
<number>
Placenta sensitivity versus placenta depth
Muscle blood volume: 25 µMol
Muscle SatO2 = 60% Placenta SatO2 = 80%
Placental sensitivity decreases with increasing placenta depth
Thickness percentiles (%) | Placenta depth (mm) | Skin thickness (mm) | Adipose tissue thickness (mm) | Muscle thickness (mm)
25 | 10 | 1 | 2 | 7
50 | 16 | 2 | 4 | 10
75 | 20 | 3 | 5 | 12
Placenta depth: 20 mm
Placenta depth: 10 mm
Placenta depth: 16 mm
Placenta depth: 10 mm

## Slide 15
<number>
Phantom calibration procedure
Solid phantom measurements with Mini CYRIL [3]
Monte carlo simulations of the solid phantom
Scale measurements to MC data
Calculate minimum detectable signal
in MC data
Calculate probability of MC data to be above measurement noise level
Emitted light source
3 cm
4 cm
5 cm
Source-detector
separation (cm)
µa=0.05 cm-1, µs = 9.3 cm-1, n = 1.37, g = 0.9
µa=0.98 cm-1, µs = 12 cm-1, n = 1.37, g = 0.9
µa=0.11 cm-1, µs = 8 cm-1, n = 1.37, g = 0.9
µa=0.1 cm-1, µs = 10 cm-1, n = 1.37, g = 0.9
Thickness (mm)
0
4
6
9
109
In-house developed broadband NIRS device
[3] Kaynezhad, P. et al., Neurophotonics 6(4), 045009–045009 (2019).

## Slide 16
<number>
Probability of acquiring a signal above the noise level
Muscle blood volume: 25 µMol
Muscle SatO2 = 60% Placenta SatO2 = 80%
Integration time 1 s
Detection probability increasing with increasing temporal binning
Thickness percentiles (%) | Placenta depth (mm) | Skin thickness (mm) | Adipose tissue thickness (mm) | Muscle thickness (mm)
25 | 10 | 1 | 2 | 7
50 | 16 | 2 | 4 | 10
75 | 20 | 3 | 5 | 12
Integration time 1 s - temporal binning 10 s
Placenta depth: 10 mm
Placenta depth: 10 mm
Placenta depth: 16 mm
Placenta depth: 20 mm
Placenta depth: 16 mm
Placenta depth: 20 mm

## Slide 17
<number>
Placenta Sensitivity calculated for 142 participants
Probability that the light has passed through the placenta
Distribution of probabilities calculated for 142 participants
<number>
5.3 %
9 %
13.4 %
Placental sensitivity increases with increasing source-detector separation

## Slide 18
<number>
Placenta Scanning probability calculated for 142 participants
Probability of acquiring a signal above the noise level and that the light has passed through the placenta - Distribution of probabilities calculated for 142 participants
<number>
4.2 %
5 %
0 %
7.6 %
0 %
0 %
Placental scanning probability increases with increasing temporal binning
5.3 %
9 %
13.4 %

## Slide 19
<number>
Discussion and Conclusion
Monte Carlo framework optimises NIRS acquisition parameters for placental monitoring
Ultrasound imaging is necessary before NIRS acquisition to evaluate placenta depth
CW-NIRS devices are more effective for placentas closer to the surface
  Placenta sensitivity is ≤ 7 % for a 2 cm placenta depth
  Placenta sensitivity is 39 % for a 1 cm placenta depth using 5 cm source-detector separation
Acquisition parameters have a major impact on placenta sensitivity
Placenta sensitivity could be enhanced with an optimized device: LED-based multi-wavelength NIRS system
Future works
Evaluate placental function with the measurement of placental tissue oxygenation

## Slide 20
Thank you for your attention
W2B.1 Luca Giannoni et al. Standardised testing procedure based on optical phantoms for validation of photonics instrumentation aimed at clinical monitoring of the placenta
Tu2A.23 Uzair Hakim et al. Optimising Placenta Oxygenation Estimation: A Monte Carlo Simulation Study
W1B.5 Frederic Lange et al. Optimizing deep tissue monitoring with diffuse optics: virtual tools for TD-NIRS and TD-SCOS development

## Slide 21
<number>
Placenta sensitivity versus melanosome volume fraction
Muscle blood volume: 25 µMol
Muscle SatO2 = 60% Placenta SatO2 = 80%
Thickness percentiles (%) | Placenta depth (mm) | Skin thickness (mm) | Adipose tissue thickness (mm) | Muscle thickness (mm)
25 | 10 | 1 | 2 | 7
Placental sensitivity increases with increasing melanosome volume fraction
Melanosome vol. fraction: 2.55 %
Placenta depth: 10 mm
Melanosome vol. fraction: 15.5 %
Melanosome vol. Fraction: 30.5 %

## Slide 22
Probability that the light has passed through the placenta
<number>
Placenta sensitivity – Melanin effect
Further investigation is needed!

## Slide 23
Probability that the signal collected by a real device is above the noise level
1) Perform measurements Mes of the multi-layered phantom with a CW NIRS device (Mini CYRIL [2])
- Source-detector separation (3 cm, 4 cm, 5 cm)
- Several integration time (1 s, 5 s, 10 s)
2) Perform MC simulations MC of the multi-layered phantom (1000 points)
3) Calculate measurement Signal to Noise Ratio:
4) Scale Measurements to MC data:
5) Calculate minimum detectable diffuse reflectance value:
<number>
Detection probability – Phantom calibration procedure
[2] Kaynezhad, P. et al., Neurophotonics 6(4), 045009–045009 (2019).

## Slide 24
Detection Probability
Probability that the signal collected is above the noise level (minimum detectable diffuse reflectance)
Detection probability PD (one-sided T-Test):
With γ a vector a 1000 elements representing the sensor noise, having a Normal distribution
<number>
Detection probability – Phantom calibration procedure

## Slide 25
<number>
Probability of acquiring a signal above the noise level
Muscle blood volume: 25 µMol
Muscle SatO2 = 60% Placenta SatO2 = 80%
Thickness percentiles (%) | Placenta depth (mm) | Skin thickness (mm) | Adipose tissue thickness (mm) | Muscle thickness (mm)
50 | 16 | 2 | 4 | 10
Integration time 1 s
Detection probability decreases with increasing source-detector separation

## Slide 26
Placenta depth: 16 mm
<number>
Probability of acquiring a signal above the noise level
Muscle blood volume: 25 µMol
Muscle SatO2 = 60% Placenta SatO2 = 80%
Integration time 1 s
Detection probability decreases with increasing placenta depth
Thickness percentiles (%) | Placenta depth (mm) | Skin thickness (mm) | Adipose tissue thickness (mm) | Muscle thickness (mm)
25 | 10 | 1 | 2 | 7
50 | 16 | 2 | 4 | 10
75 | 20 | 3 | 5 | 12
Placenta depth: 10 mm
Placenta depth: 20 mm

## Slide 27
<number>
Muscle blood volume: 25 µMol
Muscle SatO2 = 60% Placenta SatO2 = 80%
Thickness percentiles (%) | Placenta depth (mm) | Skin thickness (mm) | Adipose tissue thickness (mm) | Muscle thickness (mm)
50 | 16 | 2 | 4 | 10
Integration time 1 s
Detection probability decreases with increasing melanosome volume fraction
Probability of acquiring a signal above the noise level
Melanosome vol. fraction: 2.55 %
Melanosome vol. fraction: 30.5 %
Melanosome vol. fraction: 15.5 %

## Slide 28
<number>
Placenta sensitivity
Probability that the light has passed through the placenta
Muscle blood volume: 25 µMol
Muscle SatO2 = 60% Placenta SatO2 = 80%
Thickness percentiles (%) | Placenta depth (mm) | Skin thickness (mm) | Adipose tissue thickness (mm) | Muscle thickness (mm)
25 | 10 | 1 | 2 | 7
50 | 16 | 2 | 4 | 10
75 | 20 | 3 | 5 | 12

## Slide 29
Probability that the signal collected by a real device is above the noise level
1) Perform measurements Mes of the multi-layered phantom with a CW NIRS device (Mini CYRIL [2])
- Source-detector separation (3 cm, 4 cm, 5 cm)
- Several integration time (1 s, 5 s, 10 s)
2) Perform MC simulations MC of the multi-layered phantom (1000 points)
<number>
Detection probability – Phantom calibration procedure
Emitted light source
3 cm
4 cm
5 cm
Source-detector
separation (cm)
µa=0.05 cm-1, µs = 9.3 cm-1, n = 1.37, g = 0.9
µa=0.98 cm-1, µs = 12 cm-1, n = 1.37, g = 0.9
µa=0.11 cm-1, µs = 8 cm-1, n = 1.37, g = 0.9
µa=0.1 cm-1, µs = 10 cm-1, n = 1.37, g = 0.9
Thickness (mm)
0
4
6
9
109
[2] Kaynezhad, P. et al., Neurophotonics 6(4), 045009–045009 (2019).

## Slide 30
<number>
Detection probability – Phantom calibration procedure

## Slide 31
Estimation with standard spatially resolved spectroscopy algorithm (SRS) [7]
Hypothesis: homogeneous tissue
Simulated data (multi-distance broadband spectrum 780 nm to 890 nm)
Placental tissue oxygenation estimation
with
Placental tissue oxygenation
µa is mainly due to Hb and HbO2 contributions:
<number>
[7] Suzuki, S. Et al. in [Optical tomography and spectroscopy of tissue III ], 3597, 582–592, SPIE (1999)

## Slide 32
Placental tissue oxygenation estimation
<number>
Subject ID | Thickness percentiles (%) | Placenta depth (mm) | Skin thickness (mm) | Adipose tissue thickness (mm) | Muscle thickness (mm)
1 | 25 | 10 | 1 | 2 | 7
2 | 50 | 16 | 2 | 4 | 10
3 | 75 | 20 | 3 | 5 | 12

## Slide 33
Placental tissue oxygenation estimation
<number>